# Setup E PAT Rig Swept-Capsule Effective Sample Size Report

Setup: setup_e

Report timestamp: 20260422_095230

Source notebook: setup_E_PATrig_effective_sample_size.ipynb

# Executive summary

Setup E treats the PAT rig as a covered-probe vertical sampling interface in which powder continuously moves past a side-window non-contact NIR measurement location. The primary effective sample size is now the unique exposed optical mass represented by the initial circular 25 mm NIR footprint plus the swept capsule extension created by powder motion during acquisition. The report also retains instantaneous probed mass, swept-in incremental mass, mass-passed diagnostics, and a simple time-weighted contribution proxy for future research.

# Key parameters

Pipe diameter (cm): 10

NIR spot diameter (cm): 2.5

Base measurement depth (cm): 0.1

Measurement depth sweep (mm): 0.1 to 2.0

Bulk density scenarios (g/cm^3): 0.4, 0.5, 0.6

Base throughput (kg/h): 95

Acquisition time sweep (s): 1.0 to 5.0

Throughput sweep (kg/h): 10 to 100

Primary ESS proxy: unique exposed swept capsule mass

## Background

Setup E models a PAT (Process Analytical Technology) rig as a vertical cylindrical sampling interface fed continuously from an upstream manufacturing unit. A rotating outlet valve at the bottom controls discharge, while level sensors along the wall maintain the rig at approximately half full during operation.

The primary NIR instrument views the powder through a side optical window positioned near the lower part of the rig. In the base case used here, the measurement location is one-quarter of the rig height above the bottom outlet and the non-contact NIR spot diameter is 25 mm.

At the start of a spectral acquisition, all powder inside the circular NIR footprint contributes to the measured spectrum. During the acquisition, powder motion translates new material into the beam. Because the Setup E NIR spot is large relative to the powder advance during a 1 s acquisition, the effective sample size should include both the initial circular footprint and the additional swept-in material. The primary proxy in this notebook therefore represents the unique exposed powder as a rounded-end capsule footprint multiplied by optical penetration depth.

## Objective

Compute three Setup E sampling quantities:

instantaneous probed mass at acquisition start

unique exposed effective sample size using swept capsule geometry

time-weighted contribution proxy retained as a future research quantity

The analysis sweeps bulk density, penetration depth, acquisition time, and throughput while keeping the covered-probe PAT rig geometry and 25 mm NIR spot fixed.

## Theory

### Covered-Probe PAT Rig Geometry

The PAT rig is represented as a vertical cylindrical flow region observed through a side optical window. The operating assumption is that level control keeps the probe window buried in powder during normal operation, so probe coverage is treated as guaranteed.

### Throughput-Derived Powder Velocity

Powder motion near the covered probe is approximated as a first-order dense moving column. The local axial velocity is estimated from throughput and the full cylindrical cross-section:

where

This is a simplifying assumption. It ignores wall slip, local recirculation, outlet-valve kinematics, and nonuniform velocity profiles near the probe.

### Quantity 1: Instantaneous Probed Mass

The non-contact NIR spot is represented as a circular optical footprint:

The instantaneous probed mass at acquisition start is:

### Quantity 2: Unique Exposed Effective Sample Size

During an acquisition of length T_acq, powder advances by:

The unique footprint exposed at least once is approximated as a swept capsule: the initial circular spot plus a rectangular swept increment of width D_spot and length Delta z:

The primary effective sample size proxy is therefore:

This proxy counts unique material exposed at least once. It does not weight by residence time or optical intensity.

### Quantity 3: Time-Weighted Contribution Proxy

A measured NIR spectrum is time-integrated, so material that remains in the beam longer can contribute more strongly than material briefly entering near an edge. This notebook retains a simple mass-time proxy:

This is retained for future research only. It is not the primary effective sample size in this notebook.

### Related Diagnostic: Mass Passed the NIR Footprint

The mass passed diagnostic is:

This quantity is useful as a transport reference, but it should not be interpreted as the optical effective sample size because it does not include penetration-depth reduction.

## Interpreting Density, Velocity, and Throughput

At fixed mass throughput, the local powder velocity in this model is inversely proportional to bulk density:

For the same throughput and PAT rig cross-section, lower bulk-density powder occupies a larger volumetric flow rate and therefore moves faster past the probe. Higher bulk-density powder moves more slowly.

This has two consequences in the capsule model:

Increasing bulk density increases the instantaneous mass in the initial circular NIR footprint.

Increasing bulk density decreases the powder advance during acquisition, which reduces the swept-in incremental contribution.

Unlike a pure renewal-count proxy, density does not fully cancel in the primary Setup E effective sample size because the initial circular footprint is always included.

## Three Sampling Quantities Used in This Notebook

This notebook separates three related but distinct quantities:

Instantaneous probed mass: material inside the circular NIR footprint at acquisition start.

Unique exposed effective sample size: all material exposed at least once during acquisition, represented by the swept capsule footprint.

Time-weighted contribution proxy: a simple mass-time exposure proxy retained for future research.

The unique exposed effective sample size is the primary reported ESS in this notebook. The time-weighted proxy should not be interpreted as the primary ESS.

## Assumptions

vertical cylindrical PAT rig with top feed and bottom discharge through a rotating outlet valve

the operating window is controlled so the probe window remains continuously covered during normal operation

no explicit level-control dynamics are modeled because probe coverage is treated as guaranteed by design

the powder near the probe is treated as a dense moving column

powder velocity at the probe is approximated from throughput and full pipe cross-section

the NIR spot is circular and viewed through a side optical window

powder motion during one acquisition is represented as linear translation across the spot

the unique exposed footprint is represented as a swept capsule: initial circle plus rectangular extension

measurement depth is uniform over the full exposed capsule footprint

the capsule model counts material exposed at least once, regardless of how long it remains inside the beam

no explicit outlet-valve kinematics, turbulence, segregation, recirculation, or wall-slip correction

no detailed radiative-transfer model or optical intensity weighting is included

## Compute Sampling Metrics

Run a discrete bulk-density comparison together with a penetration-depth sweep by density, an acquisition-time sweep by density, and a throughput-acquisition-time sweep by density. The primary effective sample size is the unique exposed mass from the swept capsule footprint.

# Worked example

The baseline case reports the center density scenario, base-case penetration depth, powder advance during one acquisition, instantaneous circular mass, swept-in incremental mass, unique exposed capsule mass, the resulting primary effective sample size, the transport mass-passed diagnostic, and the future-research time-weighted proxy.

| Quantity | Value |
| --- | --- |
| bulkdensitygpercm3 | 0.5 |
| measurementdepthcm | 0.1 |
| measurementdepthmm | 1 |
| nirspotdiameter_cm | 2.5 |
| velocitymms | 6.71988 |
| acquisitiontimes | 1 |
| advancedzmmperacq | 6.71988 |
| probemassmg_instant | 245.437 |
| sweptincrementmass_mg | 83.9984 |
| uniqueexposedmass_mg | 329.435 |
| effectivesamplesize_mg | 329.435 |
| masspassedmgperacq | 1649.31 |
| timeweightedcontributionproxymg_s | 245.437 |

# Model scenarios and sweep design

The Setup E analysis combines a discrete bulk-density comparison with a penetration-depth sweep by density, an acquisition-time sweep by density, and a throughput-by-acquisition-time sweep by density. The primary response is the swept-capsule unique exposed mass.

bulk density scenarios: 0.4, 0.5, and 0.6 g/cm^3

penetration depth swept from 0.1 to 2.0 mm in 0.1 mm steps for each density

acquisition time swept from 1.0 to 5.0 s for each density

throughput swept from 10 to 100 kg/h for each density

NIR spot diameter fixed at 2.5 cm

This produces 1323 modeled rows across the density-aware sweeps.

# Selected results table

The selected table compares the three bulk-density scenarios at the common base operating condition. It separates instantaneous circular mass, swept-in incremental mass, unique exposed mass, and the final primary ESS.

| bulkdensitygpercm3 | velocitymms | acquisitiontimes | advancedzmmperacq | probemassmg_instant | sweptincrementmass_mg | uniqueexposedmass_mg | effectivesamplesize_mg | masspassedmgperacq | timeweightedcontributionproxymg_s |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0.4000 | 8.3998 | 1.0000 | 8.3998 | 196.3495 | 83.9984 | 280.3480 | 280.3480 | 1649.3056 | 196.3495 |
| 0.5000 | 6.7199 | 1.0000 | 6.7199 | 245.4369 | 83.9984 | 329.4354 | 329.4354 | 1649.3056 | 245.4369 |
| 0.6000 | 5.5999 | 1.0000 | 5.5999 | 294.5243 | 83.9984 | 378.5228 | 378.5228 | 1649.3056 | 294.5243 |

# Figures

The figures are ordered from the geometric decomposition of effective sample size to penetration-depth trends, powder-advance behavior, operating maps, the future-research time-weighted proxy, and the capsule schematic.

Figure 1. Instantaneous circular mass, swept-in incremental mass, and unique exposed ESS versus penetration depth for the base density case.

Figure 2. Swept-capsule effective sample size versus penetration depth for bulk densities of 0.4, 0.5, and 0.6 g/cm^3.

Figure 3. Powder advance during acquisition versus acquisition time for the three bulk-density scenarios.

Figure 4. Swept-capsule effective sample size versus acquisition time for the three bulk-density scenarios.

Figure 5. Swept-in incremental mass versus acquisition time for the three bulk-density scenarios.

Figure 6. Throughput-acquisition heatmap for swept-capsule effective sample size at 0.4 g/cm^3 bulk density.

Figure 7. Throughput-acquisition heatmap for swept-capsule effective sample size at 0.5 g/cm^3 bulk density.

Figure 8. Throughput-acquisition heatmap for swept-capsule effective sample size at 0.6 g/cm^3 bulk density.

Figure 9. Time-weighted contribution proxy versus acquisition time, retained as a future-research quantity rather than the primary ESS.

Figure 10. Schematic of the swept capsule geometry used for the primary unique exposed effective sample size proxy.

## Results Interpretation

The primary effective sample size now includes all material exposed at least once during acquisition.

The base case contains the full initial circular spot plus an additional swept-in strip generated by powder motion during the 1 s acquisition.

Because the 25 mm spot is large relative to the base-case powder advance, the instantaneous component remains a major part of the effective sample size.

Increasing acquisition time increases effective sample size by elongating the exposed capsule footprint.

Increasing penetration depth increases instantaneous mass, swept-in incremental mass, and effective sample size linearly.

Lower bulk density increases powder velocity and the swept increment at fixed throughput, while higher bulk density increases the mass in the initial circular footprint.

The time-weighted contribution proxy is retained only as a future route toward residence-time and optical-weighting analysis.

## Practical Implications

Large NIR spots at low powder velocity should not use a refresh-fraction-only effective sample size proxy.

The starting circular footprint can dominate the effective sample size when powder advance is small relative to spot diameter.

Acquisition time elongates the exposed footprint rather than simply replacing a fraction of the original spot.

For Setup E, the swept capsule model better matches the all-unique-exposed-material logic used elsewhere in the project.

Throughput directly controls powder advance and therefore the incremental swept contribution, but the initial circular contribution remains present for every spectrum.

## Limitations

probe coverage is assumed guaranteed and is not stress-tested against abnormal operating states

the rotating outlet valve is not modeled dynamically

the local velocity near the probe is approximated from throughput and full pipe cross-section

straight-line translation of powder past the spot is assumed

velocity is assumed uniform over the NIR spot footprint

penetration depth is uniform over the full capsule footprint

residence-time differences across the swept capsule are not included in the primary ESS

radial and tangential optical intensity gradients are ignored

penetration depth is swept as an effective parameter, not derived from a wavelength-dependent optical model

bulk density is treated as spatially uniform and fixed within each scenario

no particle-level sampling statistics, recirculation, stagnant wall-layer, or repeated-exposure correction is included

# Appendices

## Next Logical Extensions

replace the top-hat capsule geometry with nonuniform optical intensity weighting

model local velocity distribution near the side window

include residence-time weighting over the swept capsule

couple penetration depth to density and wavelength

add uncertainty bands for penetration depth, bulk density, and local velocity

validate the capsule effective sample size proxy against tracer experiments or PAT spectral-response data

# Exported artifacts

CSV files

results_full.csv

summary_table.csv

penetration_depth_table.csv

acquisition_time_table.csv

input_summary.csv

Figure files

plot_01_instant_swept_unique_vs_penetration_depth.png

plot_02_effective_sample_size_vs_penetration_depth_by_density.png

plot_03_powder_advance_vs_acquisition_time.png

plot_04_effective_sample_size_vs_acquisition_time_by_density.png

plot_05_swept_increment_mass_vs_acquisition_time_by_density.png

plot_06_heatmap_effective_sample_size_density_04.png

plot_07_heatmap_effective_sample_size_density_05.png

plot_08_heatmap_effective_sample_size_density_06.png

plot_09_time_weighted_proxy_vs_acquisition_time_by_density.png

schematic_swept_capsule_geometry.png