# Setup E PAT Rig Effective Sample Size Report

Setup: setup_e

Report timestamp: 20260421_090917

Source notebook: setup_E_PATrig_effective_sample_size.ipynb

# Executive summary

Setup E treats the PAT rig as a level-controlled vertical sampling interface in which a half-filled dense powder column moves past a side-window non-contact NIR measurement location. The refactored study separates three sensitivities: penetration depth, acquisition time, and throughput-acquisition operating conditions, while keeping the 25 mm spot geometry fixed so that illuminated mass, refreshed new mass, and transported mass can be compared mechanistically.

# Key parameters

Pipe diameter (cm): 10

Total rig height (cm): 80

Maintained fill fraction (-): 0.5

Modeled filled height (cm): 40

Probe height from bottom (cm): 20

NIR spot diameter (cm): 2.5

Base measurement depth (cm): 0.1

Measurement depth sweep (mm): 0.1 to 2.0

Bulk density (g/cm^3): 0.3

Base throughput (kg/h): 95

Acquisition time sweep (s): 1.0 to 5.0

Throughput sweep (kg/h): 10 to 100

Axial length model: spot_diameter

## Background

Setup E models a PAT (Process Analytical Technology) rig as a vertical cylindrical sampling interface fed continuously from an upstream manufacturing unit. A rotating outlet valve at the bottom controls discharge, while level sensors along the wall maintain the rig at approximately half full during operation.

The primary NIR instrument views the powder through a side optical window positioned near the lower part of the rig. In the base case used here, the measurement location is placed one-quarter of the rig height above the bottom outlet and the non-contact NIR spot diameter is 25 mm.

Three distinct mass metrics are relevant:

Instantaneous probed mass (m_inst): mass simultaneously illuminated at any single instant.

New mass presented per acquisition (m_new): unique material refreshed across the NIR spot during one acquisition time, bounded by the instantaneous mass.

Mass passed probe footprint (m_pass): total mass that crosses the probe footprint during acquisition and is not bounded by the instantaneous mass.

This notebook treats the occupied lower section of the PAT rig as a first-order dense moving powder column. It uses the maintained fill level to determine whether the probe is covered by powder and then uses throughput-driven axial motion to estimate renewal at the probe position.

## Objective

Compute instantaneous probed mass, new mass presented per acquisition, and mass passed the NIR spot footprint for a level-controlled PAT rig configuration. Use a half-filled base case with the measurement window located one-quarter of the rig height from the bottom, sweep acquisition time and throughput, and export results using the standard reporting workflow.

## Theory

### Level-controlled PAT rig geometry

The PAT rig is modeled as a vertical cylinder of total height Hrig and internal diameter Dpipe. The level-control system maintains an approximately constant filled fraction f_fill, so the modeled filled height is:

The NIR measurement location is placed at height z_probe above the outlet. A simple coverage indicator is then used:

C_probe = begincases1, & z_probe le H_fill \\ 0, & z_probe > H_fillendcases

In the base case, the probe sits below the maintained fill level, so the optical window is treated as continuously covered by powder.

### Throughput-driven axial velocity

Within the occupied lower section of the PAT rig, powder motion near the probe is approximated as a first-order dense moving column. The axial velocity is estimated from throughput and the full cylindrical cross-section:

where

This is still a simplifying assumption: it ignores detailed wall slip, outlet-valve kinematics, and any nonuniform velocity profile near the probe.

### Instantaneous probed mass

The non-contact NIR spot is represented as a circular optical footprint with area:

The instantaneous probed mass, if the probe is covered, is:

### Axial advance and refresh fraction

Axial advance during one acquisition is:

The axial sampling length is represented by a proxy length L_axial. In the base case this is taken as the NIR spot diameter. The refresh fraction is then:

### New mass and mass passed

New mass presented per acquisition is:

Mass passed the probe footprint during acquisition is:

Symbols and units:

dotm: mass flow rate [g/s]

rho_bulk: bulk density [g/cm^3]

A_pipe: pipe cross-section area [cm^2]

H_rig: total rig height [cm]

f_fill: maintained filled fraction [-]

H_fill: modeled filled height [cm]

z_probe: probe height above the outlet [cm]

C_probe: probe coverage indicator [-]

A_spot: NIR spot area [cm^2]

D_spot: NIR spot diameter [cm]

d_pen: measurement (penetration) depth [cm]

T_acq: spectral acquisition time [s]

L_axial: axial sampling length proxy [cm]

## Assumptions

vertical cylindrical PAT rig with top feed and bottom discharge through a rotating outlet valve

level sensors maintain an approximately constant filled fraction during operation

the occupied lower section is treated as a dense moving powder column near the probe

powder velocity at the probe is approximated from throughput and full pipe cross-section

the NIR spot is circular and viewed through a side optical window

the probe is treated as covered whenever its height is below the modeled fill height

measurement depth is small relative to pipe diameter

axial sampling length is approximated by the NIR spot diameter unless overridden

no explicit outlet-valve kinematics, turbulence, segregation, or wall-slip correction

no detailed radiative transfer model yet

## Compute Effective Sample Size Metrics

Run a 1D penetration-depth sweep, a 1D acquisition-time sweep, and a 2D throughput x acquisition-time sweep while keeping the half-filled PAT rig geometry and 25 mm NIR spot fixed.

# Worked example

The baseline case below reports the maintained fill height, probe position relative to the fill level, the base-case penetration depth, the instantaneous mass at the probe, the refreshed new mass during one acquisition, the total mass passed the NIR spot footprint, and the resulting refresh fraction.

| Quantity | Value |
| --- | --- |
| filledheightcm | 40 |
| probeheightfrombottomcm | 20 |
| probesubmergencecm | 20 |
| probeiscovered | 1 |
| measurementdepthcm | 0.1 |
| velocitymms | 11.1998 |
| acquisitiontimes | 1 |
| refreshfractionper_acq | 0.447992 |
| probemassmg_instant | 147.262 |
| newmassmgperacq | 65.9722 |
| masspassedmgperacq | 1649.31 |
| effectivesamplesize_mg | 65.9722 |

# Model scenarios and sweep design

The Setup E analysis combines a one-dimensional penetration-depth sweep, a one-dimensional acquisition-time sweep, and a two-dimensional throughput-by-acquisition-time sweep. This separates optical-depth sensitivity from renewal-time sensitivity while holding the PAT rig geometry and 25 mm spot size fixed.

penetration depth swept from 0.1 to 2.0 mm in 0.1 mm steps

acquisition time swept from 1.0 to 5.0 s

throughput swept from 10 to 100 kg/h

maintained fill fraction fixed at 0.50

probe height fixed at 20.0 cm from the bottom

NIR spot diameter fixed at 2.5 cm

axial sampling length model: spot_diameter

This produces 440 modeled rows across the 1D and 2D sweeps.

# Selected results table

The selected table below samples representative acquisition times from the revised 1-5 s sweep so the reader can compare fill condition, base penetration depth, instantaneous illuminated mass, refreshed new mass, passed mass, refresh fraction, and the resulting effective sample-size proxy side by side.

| acquisitiontimes | measurementdepthcm | velocitymms | filledheightcm | probeheightfrombottomcm | probeiscovered | probemassmg_instant | newmassmgperacq | masspassedmgperacq | refreshfractionper_acq | effectivesamplesize_mg |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1.0000 | 0.1000 | 11.1998 | 40.0000 | 20.0000 | True | 147.2622 | 65.9722 | 1649.3056 | 0.4480 | 65.9722 |
| 2.0526 | 0.1000 | 11.1998 | 40.0000 | 20.0000 | True | 147.2622 | 135.4167 | 3385.4167 | 0.9196 | 135.4167 |
| 3.1053 | 0.1000 | 11.1998 | 40.0000 | 20.0000 | True | 147.2622 | 147.2622 | 5121.5278 | 1.0000 | 147.2622 |
| 5.0000 | 0.1000 | 11.1998 | 40.0000 | 20.0000 | True | 147.2622 | 147.2622 | 8246.5278 | 1.0000 | 147.2622 |

# Figures

The figures are ordered from one-dimensional depth and time mechanism checks to two-dimensional operating maps: effective sample size versus penetration depth, instantaneous probed mass versus penetration depth, effective sample size versus acquisition time, refresh fraction versus acquisition time, then the throughput-acquisition heatmaps.

Figure 1. Effective sample size versus penetration depth for the half-filled PAT rig base case.

Figure 2. Instantaneous probed mass versus penetration depth for the 25 mm NIR spot.

Figure 3. Effective sample size versus acquisition time at the base penetration depth.

Figure 4. Refresh fraction versus acquisition time for the 25 mm NIR spot.

Figure 5. Throughput-acquisition heatmap for the effective sample-size proxy.

Figure 6. Throughput-acquisition heatmap for mass passed the NIR spot footprint.

## Results Interpretation

in the half-filled base case, the modeled fill height remains above the measurement location, so the NIR window is treated as continuously covered

increasing penetration depth raises both instantaneous illuminated mass and the effective-sample-size ceiling

increasing acquisition time raises refresh fraction and effective sample size until transport renewal approaches saturation

throughput controls powder velocity and therefore the operating envelope seen in the 2D heatmap

the 25 mm NIR spot is fixed across all sweeps, so the trends isolate depth, time, and throughput effects rather than changes in spot geometry

## Practical Implications

penetration depth uncertainty matters directly because it changes the illuminated mass and the resulting effective sampled mass proxy

acquisition time selection still trades off spectral quality against how much unique material is refreshed through the side-window NIR measurement

throughput directly controls powder velocity and therefore refresh rate near the measurement location

maintaining the rig at approximately half full provides a simple coverage condition while the 25 mm spot remains the fixed instrument geometry assumption

## Limitations

the half-filled condition is represented by a fixed fill fraction, not by an explicit level-control simulation

the rotating outlet valve is not modeled dynamically

the local velocity near the probe is approximated from throughput and full pipe cross-section

no turbulence, wall-slip, or recirculation model is included

penetration depth is swept as an effective parameter, not derived from a wavelength-dependent optical model

bulk density is assumed uniform and constant

no particle-level sampling statistics or optical weighting correction are included

# Appendices

## Next Logical Extensions

dynamic level-control and rotating-valve discharge model

explicit dependence of local velocity on valve operation and bed holdup

sensitivity analysis for probe height and maintained fill fraction

uncertainty bands for measurement depth and bulk density

non-uniform optical weighting functions for the 25 mm NIR spot

# Exported artifacts

CSV files

results_full.csv

summary_table.csv

penetration_depth_table.csv

input_summary.csv

Figure files

plot_01_effective_sample_size_vs_penetration_depth.png

plot_02_instantaneous_probed_mass_vs_penetration_depth.png

plot_03_effective_sample_size_vs_acquisition_time.png

plot_04_refresh_fraction_vs_acquisition_time.png

plot_05_heatmap_effective_sample_size_throughput_vs_acquisition.png

plot_06_heatmap_mass_passed_throughput_vs_acquisition.png