# Setup D Effective Sample Size Report

# Run Metadata

Setup name: setup_d

Run timestamp: 20260317_112555

Notebook filename: setup_D_pipe_effective_sample_size.ipynb

Run directory: C:\temp files\EFS\reports\setup_d\20260317_112555

# Notebook Narrative

# Setup D: inclined pipe effective sample size proxy

## 1. Title and purpose

This notebook develops a first-order engineering model for Setup D, an inclined powder pipe with an NIR diffuse reflectance probe mounted in the lower part of the pipe. The objective is to estimate a flow-based effective sample size proxy during one spectrum acquisition.

The result is best interpreted as a gross refreshed optical mass / effective sample size proxy. It is not a full radiative-transfer model, not a particulate sampling-theory model, and not a validated estimate of independent sample mass.

## 2. Scientific background

### NIR diffuse reflectance analysis of powders

In diffuse reflectance NIR measurements, light enters a powder bed, interacts with particles over a limited depth, and a fraction of the scattered light returns to the detector. The measured spectrum therefore reflects an optically weighted near-surface volume rather than the entire bulk stream.

### Why use a simple optical interaction approximation?

For engineering comparisons, it is often useful to approximate the optical interaction region with a simple geometric volume. Here, the optical sampling region is approximated as a cylinder with:

cross-sectional area equal to the probe window footprint

depth equal to an assumed light penetration depth

This simplification ignores depth-dependent photon weighting, scattering anisotropy, particle packing effects, and spectral absorption details, but it provides a practical way to compare conditions.

### Why moving powder increases the sampled mass during one acquisition

If powder continuously moves past the probe during one acquisition, the optical interaction zone is refreshed. Faster renewal means that more powder volume can pass through the probe footprint during the acquisition period. That does not guarantee statistical independence of all material seen by the instrument, but it increases the gross refreshed optical mass available to contribute to the measurement.

### Why Setup D differs from other geometries

Setup D differs from a static powder bed because the material in front of the probe is continuously replaced. It also differs from a mechanically refreshed geometry such as the spiderwheel because the refresh mechanism is not imposed by discrete compartments or controlled mechanical presentation. Instead, it is driven by gravity-assisted powder transport through a tilted pipe.

### Why a superficial velocity model is useful

A practical first approximation is to connect throughput, bulk density, and pipe cross-sectional area to a superficial bulk velocity. This gives a simple estimate of how quickly powder is renewed in front of the probe. In reality, the local flow at the lower wall of a tilted pipe may differ substantially from the cross-sectional average, but the superficial velocity remains a useful baseline for sensitivity analysis and engineering intuition.

## 3. Setup D device description

Setup D is modeled as follows:

pipe internal diameter = 100 mm

pipe axis tilted 15 degrees from vertical

powder enters from above and moves downward toward the outlet

due to the tilt, powder is expected to slide preferentially along the lower side of the pipe

the NIR probe is mounted in the lower part of the pipe

the probe window is assumed to be continuously covered by powder

the effective sample size proxy therefore depends primarily on powder renewal in front of the probe

This notebook uses a first-order flow approximation. The 15 degree tilt is included as part of the physical setup description and qualitative interpretation, but the base model still estimates renewal from the total throughput, bulk density, and full pipe cross-sectional area.

## 4. Model assumptions

The base model uses the following assumptions:

the optical sampling volume is approximated as a cylinder: probe window area x penetration depth

the probe window is fully covered by powder at all times

powder renewal at the probe is approximated by a bulk superficial velocity derived from throughput, density, and pipe cross-sectional area

bulk density is constant and spatially uniform

no explicit blocking events occur

no recirculation correction is included

local wall velocity may differ from average superficial velocity, but that difference is ignored in the base model

the output is a flow-based effective sample size proxy, not a validated independent sample mass

These assumptions intentionally favor transparency and reuse over physical completeness.

## 5. Mathematical model

Internal units used throughout the notebook are:

geometry in cm

time in s

density in g/cm^3

throughput entered in kg/h and converted internally to g/s

penetration depth entered in mm and converted internally to cm

final reported effective sample size proxy in mg

Pipe radius:

Pipe cross-sectional area:

Mass flow conversion:

Superficial velocity:

Probe footprint area:

Optical volume per static slice:

Optical mass per static slice:

Crossing time of one probe diameter:

Number of refreshes during acquisition:

Effective sample size proxy:

Equivalent transported-volume view:

Under the simplified assumptions, these forms are equivalent because the probe footprint is refreshed by translation of material past the window. The 15 degree tilt remains part of the physical setup description, but the base model still uses the cross-sectional superficial velocity as the first-order renewal velocity approximation.

## 9. Single-case worked example

The worked example uses:

throughput = 70 kg/h

density = 0.5 g/cm^3

pipe diameter = 100 mm

pipe tilt = 15 degrees from vertical

probe diameter = 5.0 mm

penetration depth = 1.0 mm

acquisition time = 1.0 s

This example should be interpreted as a first-order estimate. In a real tilted pipe, the local velocity at the lower wall may differ from the superficial velocity because the flowing layer can be asymmetric and concentrated near the lower side.

## 10. Parameter sweep across throughput, density, and penetration depth

The full sweep spans:

throughput from 40 to 100 kg/h

density = 0.4, 0.5, 0.6 g/cm^3

penetration depth from 0.1 to 2.0 mm in 0.1 mm steps

probe diameter fixed at 5 mm

acquisition time fixed at 1.0 s unless changed in the parameter cell

## 11. Results tables

## 12. Visualization

## 13. Interpretation of trends

Several clear trends appear in this first-order model.

Increasing density increases the optical mass per static slice linearly because mopt = rho Aprobe d_pen.

Increasing penetration depth increases the optical mass per static slice linearly because the modeled optical volume grows linearly with depth.

Increasing throughput increases the superficial velocity linearly for fixed density and pipe area.

Increasing superficial velocity decreases the crossing time and therefore increases the refresh count during one acquisition.

The result should be interpreted as a flow-based refreshed optical mass proxy, not necessarily an independent sampling mass.

### Density cancellation under the simplified assumptions

Starting from:

and substituting:

gives:

So, in this specific simplified model, density cancels completely for the effective sample size proxy when throughput is fixed in mass units and the whole pipe cross-section is assumed to participate in the superficial flow. Physically, a higher density increases the static optical mass seen at one instant, but it simultaneously reduces the superficial velocity in the same proportion. The larger mass per refresh and the smaller number of refreshes offset each other.

This complete cancellation is a direct consequence of the model structure. In a more realistic tilted-pipe model, density could re-enter through changes in flowing layer thickness, local wall velocity, consolidation, slip, or coverage dynamics.

## 14. Model limitations

Important limitations of this notebook are:

the cylindrical optical volume is only an approximation

there is no radiative-transfer model and no photon weighting with depth

superficial velocity may differ from the true local wall velocity at the probe location

the 15 degree tilt is represented only qualitatively in the base model

there is no correction for asymmetric sliding-layer thickness

there is no correction for stagnant regions or wall slip

there is no particle size or particle count model

there is no recirculation correction

the model is not linked directly to analytical variance or Gy sampling theory

Useful future refinements could include:

an effective moving-layer thickness

a local wall velocity factor

a filling fraction or flowing cross-sectional area correction

a residence-time or particle revisit model

## 15. Conclusions

This notebook estimates a first-order flow-based effective sample size proxy for Setup D, expressed as refreshed optical mass during one spectrum acquisition.

Within this model, the strongest explicit drivers are:

throughput, which sets superficial velocity and refresh rate

penetration depth, which scales the optical mass per refresh

probe and pipe geometry, which determine the footprint and renewal scale

The notebook is useful for sensitivity analysis, quick scenario comparison, and engineering intuition during development. A more advanced model could incorporate local lower-wall flow kinematics, effective moving-layer thickness, nonuniform filling, revisit behavior, and optical depth weighting.

## Export and Reporting

This export section writes timestamped CSV, PNG, JSON, and Word outputs into the setup-specific reports/ folder. python-docx and nbformat are required for Word export with Markdown rendering.

# Purpose

Setup D estimates a first-order flow-based effective sample size proxy in a tilted pipe by combining optical mass per refresh with a refresh rate derived from superficial velocity.

# Assumptions

Superficial velocity is used as a first-order flow approximation, density is uniform within each scenario, the optical volume is cylindrical, and the tilt effect is captured only through the stated simplified geometry.

# Parameter Summary

Pipe diameter (mm): 100

Pipe tilt from vertical (deg): 15

Probe diameter (mm): 5

Acquisition time (s): 1

Throughput range (kg/h): 40 to 100

Density scenarios (g/cm^3): 0.4, 0.5, 0.6

Penetration depth range (mm): 0.1 to 2.0

# Exported CSV Files

results_full.csv

summary_table.csv

velocity_summary.csv

Detailed tables are saved separately as CSV files in the run folder.

# Figures

plot_01_proxy_vs_penetration_depth.png

plot_02_proxy_vs_throughput.png

plot_03_heatmap_density_05.png

plot_04_superficial_velocity_vs_throughput.png

plot_05_optical_mass_per_refresh.png

plot_06_contour_density_06.png

plot_07_passed_mass_vs_throughput.png

# Closing Note

This report captures the notebook narrative, summary metadata, and exported artifacts for this run.