# Setup C Effective Sample Size Report

# Run Metadata

Setup name: setup_c

Run timestamp: 20260318_122252

Notebook filename: setup_C_spiderwheel_effective_sample_size.ipynb

Run directory: C:\Users\Erik\Documents\Projects\EFS\reports\setup_c\20260318_122252

# Notebook Narrative

# Setup C -- Spiderwheel-Based Refreshed Optical Mass Proxy

First-order mechanistic approximation for wheel-driven powder renewal and binary probe blocking during one NIR spectral acquisition.

## Background

Setup C represents a spiderwheel-driven powder presentation geometry for NIR diffuse reflectance measurement. As in the other notebooks in this series, the optical interaction region is treated as a shallow near-surface region in front of the probe window, and the optically sampled mass is approximated from probe footprint, assumed penetration depth, and bulk density.

The key physical difference from Setup B is that Setup B treats each spoon presentation as a static optical event, whereas Setup C uses wheel motion to actively renew powder at the probe location during a single acquisition. The spiderwheel therefore introduces two competing effects: it can increase the rate at which newly presented powder passes through the probe region, but the fingers can also intermittently block the probe and reduce the usable observation time.

This notebook is therefore positioned as a first-order mechanistic approximation and a deterministic renewal model. It is intended to quantify refreshed optical mass trends and blocking trends, not to claim a validated statistical effective sample size or a full radiative-transfer description of the measurement.

## Objective

This notebook estimates, for Setup C:

optical sampled mass per refresh event

blocked and unblocked acquisition time during one spectrum

number of refresh opportunities during one spectrum

refreshed optical mass proxy per spectrum

sensitivity of the proxy to penetration depth, bulk density, and wheel RPM

## Theory

Wheel kinematics at the probe position:

First-order optical footprint and optically sampled mass per refresh event:

Characteristic renewal time scale:

Effective unblocked acquisition time:

Deterministic renewal count:

Refreshed optical mass proxy during one acquisition:

In the current implementation, blocked time is estimated from a simple overlap-angle approximation in which finger width and probe width are converted into an angular overlap at the probe radius. This produces a deterministic blocked-time estimate for the acquisition window under the binary blocking approximation.

Symbol definitions and units:

RPM: wheel rotation rate [rev min^-1]

omega: angular speed [rad s^-1]

r_probe: probe radial position from wheel center [cm]

v: tangential speed at the probe radius [cm s^-1]

D_probe: probe window diameter [cm] in calculations, reported as [mm] in engineering tables and labels

A_probe: probe optical footprint area [cm^2]

d_pen: assumed optical penetration depth [cm] in calculations, reported as [mm] in engineering tables and labels

V_opt: optically sampled volume per refresh event [cm^3]

rho_bulk: bulk density [g cm^-3]

m_opt: optically sampled mass per refresh event [g]

t_acq: total spectral acquisition time [s]

t_blocked: time during which the probe is treated as blocked by a spiderwheel finger [s]

t_eff: effective unblocked acquisition time [s]

t_cross: characteristic crossing time for one probe diameter of tangential motion [s]

N_refresh: deterministic renewal count or renewal-count proxy during one acquisition [-]

m_eff,proxy: refreshed optical mass proxy during one acquisition [g]

The code uses mixed practical engineering units for convenience: geometry is entered mainly in cm and mm, density in g/cm^3, time in s, and reported masses are often shown in mg. The calculations remain internally consistent because the unit conversions are handled explicitly in the code.

## Interpretation of the Model

In this notebook, a refresh means the passage of newly presented powder across the probe region due to wheel motion. It is best interpreted as a renewal opportunity for material presented to the optical window, not as proof of a statistically independent new sample.

Tangential speed is evaluated at the probe radius because the local renewal rate should be tied to the wheel speed where the probe actually views the material. The crossing time t_cross is then interpreted as a characteristic time scale for a wheel-driven material segment to traverse one probe width.

Blocked time is subtracted because the model uses a binary blocking approximation: when a spiderwheel finger overlaps the probe region, the probe is treated as having no usable optical access to powder during that interval. The resulting t_eff is therefore an effective observation time rather than the full instrument acquisition time.

Nrefresh should therefore be interpreted as a deterministic renewal count, or renewal-count proxy, rather than a demonstrated count of independent samples. Powder recirculation, revisit probability, local residence-time memory, and correlation between successive refresh events are not modeled. For the same reason, meff,proxy is a refreshed optical mass proxy, not yet a validated statistical effective sample size and not a direct estimate of analytical variance.

Despite these simplifications, the model is still useful as a mechanistic engineering tool because it shows how geometry, density, penetration depth, RPM, and blocking assumptions influence the amount of powder that is plausibly renewed in front of the probe during one spectrum.

## Assumptions

circular probe window with area represented by A_probe

first-order cylindrical optical sampling volume approximation, Vopt = Aprobe d_pen

constant bulk density over the optically sampled region

constant assumed penetration depth for each simulation condition

local renewal rate represented by tangential velocity at the probe radius

binary blocking behavior when a spiderwheel finger overlaps the probe region

no partial transmission, partial attenuation, or partial exposure through a finger

no revisit probability or recirculation correction for previously presented powder

no depth-wise optical weighting within the assumed sampling volume

no stochastic treatment of particle independence or correlation between refresh events

no explicit radiative-transfer solution

## Input Parameters

The baseline geometric, optical, and operating parameters are defined below. Engineering units are retained in variable names, plot labels, and summary tables for readability, while the calculations remain internally consistent in centimeter-gram-second style units.

## Imports

## Helper Functions

## Single-Case Worked Example

The following worked example evaluates one representative operating condition so that the intermediate quantities can be read directly before the full parameter sweep is run. The sequence is intentionally mechanistic: first the optically sampled mass per refresh event is estimated, then blocking is estimated, then the effective unblocked acquisition time is obtained, and only then is the refreshed optical mass proxy calculated through the deterministic renewal count.

## Parameter Sweep / Simulation

The simulation sweeps bulk density, wheel RPM, and penetration depth while holding the remaining geometry fixed. Each row in the output table represents one deterministic operating condition for the current first-order model.

## Interpretation of Simulation Table

Each row combines one density, one RPM, and one penetration depth. The most important columns are:

opticalmassperrefreshmg: the first-order optically sampled mass associated with one renewal opportunity at the probe

blockedtimes: the portion of the acquisition treated as unavailable because a spiderwheel finger blocks the probe region

n_refresh: the deterministic renewal count, or renewal-count proxy, during the unblocked part of the acquisition

effectivesamplemassproxymg: the refreshed optical mass proxy accumulated over one spectrum

This table should therefore be read as a mechanistic comparison table. It does not demonstrate that the listed renewal opportunities are statistically independent, and it does not by itself establish a validated independent effective sample mass.

## Plot 1: Refreshed Optical Mass Proxy vs Penetration Depth

This plot isolates how the refreshed optical mass proxy changes with penetration depth at a representative wheel speed of 20 RPM. In the present first-order volume approximation, penetration depth scales the optically sampled volume and therefore the optically sampled mass per refresh event linearly for fixed density.

Comparing the density curves shows how increased bulk density multiplies that per-refresh optical mass contribution without changing the blocking geometry itself.

## Plot 2: Refreshed Optical Mass Proxy vs RPM

This plot isolates the operating-speed effect at a representative penetration depth of 1.0 mm. Increasing RPM increases tangential speed at the probe radius and therefore decreases the characteristic crossing time, which tends to increase the deterministic number of renewal opportunities.

That increase should not be interpreted as a direct increase in independent sample count. The plotted result is a first-order balance between faster wheel-driven renewal and the loss of effective observation time predicted by the blocking approximation.

## Plot 3: Heatmap of Refreshed Optical Mass Proxy

This heatmap presents the coupled effect of RPM and penetration depth at a representative bulk density of 0.5 g/cm^3. It is useful for identifying broad operating regions rather than a single one-dimensional trend.

Because both deterministic renewal count and optically sampled mass per refresh event contribute to the final proxy, the heatmap makes it easier to see where different RPM-depth combinations produce similar refreshed optical mass proxy values.

## Plot 4: Blocked Time Fraction vs RPM

This plot separates the blocking model from the optical-mass model. It shows how much of the acquisition is estimated to be unavailable because the spiderwheel fingers overlap the probe region under the binary blocking approximation.

The blocked fraction matters because it reduces effective observation time even in cases where wheel speed increases the rate of renewal opportunities.

## Plot 5: Optical Mass per Refresh vs Penetration Depth

This plot shows the optically sampled mass per refresh event before the renewal-count multiplier is applied. It therefore isolates the optical-volume approximation from the dynamic renewal approximation.

Reading Plot 5 together with Plot 1 helps distinguish two quantities that should not be conflated: mopt is the per-refresh optically sampled mass, whereas meff,proxy is the refreshed optical mass proxy accumulated over one acquisition.

## Plot 6: Contour Map of Refreshed Optical Mass Proxy

This contour plot provides the same coupled RPM-depth view as the heatmap, but for the higher-density case of 0.6 g/cm^3. It is useful for visualizing smooth operating envelopes and approximate iso-proxy contours.

The figure remains a first-order mechanistic map. It is useful for design sensitivity screening, but it should not be interpreted as a validated surface of statistically independent effective sample size.

## Results Interpretation

increasing bulk density increases the optically sampled mass per refresh event linearly in this simplified volume model and therefore increases the refreshed optical mass proxy proportionally when the renewal count is unchanged

increasing penetration depth increases the assumed optically sampled volume linearly and therefore increases the per-refresh optical mass linearly under the current approximation

increasing RPM decreases the characteristic crossing time and therefore increases the deterministic number of renewal opportunities, all else being equal

blocking reduces effective observation time and therefore reduces the refreshed optical mass proxy relative to an unblocked case

these outputs are mechanistically useful for engineering comparison, but they do not by themselves establish statistical independence of refresh events or a validated independent effective sample size

## Limitations

partial blocking, partial transmission, and partial attenuation are not modeled; blocking is treated as binary

recirculation, revisit probability, and repeated presentation of the same powder increment are not modeled

particle-scale heterogeneity, segregation, and particle-count statistics are outside the present approximation

depth-dependent photon weighting is not represented; the optical sampling region is treated with a uniform first-order volume approximation

residence-time distributions and local mixing memory near the probe are not modeled

stochastic independence of refresh events is not modeled, so renewal count is not equivalent to validated independent sample count

the model does not yet convert refreshed optical mass proxy into analytical variance, prediction uncertainty, or a validated statistical effective sample size

## Next Logical Extensions

replace binary blocking with a partial-blocking or partial-exposure model

include revisit probability or explicit recirculation of powder around the wheel

resolve the acquisition in time or phase so that exposure depends on wheel angle during the spectrum

introduce optical weighting with depth instead of a uniform cylindrical volume approximation

estimate particle count as well as mass within the refreshed optical region

apply a correlation correction between refresh events before interpreting the result as effective sample mass

link refreshed optical mass proxy to analytical variance or prediction precision

## Export and Reporting

This export section writes timestamped CSV, PNG, JSON, and Word outputs into the setup-specific reports/ folder. python-docx and nbformat are required for Word export with Markdown rendering.

# Purpose

Setup C estimates a first-order refreshed optical mass proxy for a spiderwheel presentation by combining optical mass per refresh with a deterministic renewal count derived from wheel speed and blocked acquisition time.

# Assumptions

Uniform bulk density, cylindrical optical sampling volume, binary full-blocking simplification for wheel fingers, and deterministic refresh counting without revisit, recirculation, partial-blocking, or correlation corrections.

# Parameter Summary

Finger count: 4

Wheel radius (cm): 3

Probe radius (cm): 1.8

Finger width (cm): 0.3

Probe diameter (mm): 5

Acquisition time (s): 1

Density scenarios (g/cm^3): 0.4, 0.5, 0.6

RPM range: 10 to 35

Penetration depth range (mm): 0.1 to 2.0

# Exported CSV Files

results_full.csv

summary_table.csv

blocking_summary.csv

Detailed tables are saved separately as CSV files in the run folder.

# Figures

plot_01_proxy_vs_penetration_depth.png

plot_02_proxy_vs_rpm.png

plot_03_heatmap_density_05.png

plot_04_blocked_time_fraction_vs_rpm.png

plot_05_optical_mass_per_refresh.png

plot_06_contour_density_06.png

# Closing Note

This report captures the notebook narrative, summary metadata, and exported artifacts for this run.